## 2026 -27 WORKING VERSION
| TMDA Report Number | Date Received | Device Brand Name | Device Common Name | Size | Batch/Lot/ Serial number | Device Type | Manufacturing Date | Expiry Date | Name and physical Address of Manufacturer | Manufacturing country | Name of the Supplier ( If applicable) | Adverse Event(s)/ Incident(s) desription | Date of Onset of Event(s)/ Incident(s) | Date of the report | Place/ Location of Event (s)/ Incident(s) | Region | Type of Report | Reporter details (Name / contact Information) | Event Seriousnes (Yes/No) | Preferred Term - Device Component Level 1: | Preferred Term of -Device Component Level 2: | Preferred Term of -Device Component Level 3: | Device Component IMDRF Codes # | Preferred Term- Device Problem Level 1: | Preferred Term - Device Problem Level 2: | Preferred Term - Device Problem Level 3: | Device Problem IMDRF Codes # | Preferred Term - Clinical Sign Level 1: | Preferred Term - Clinical Sign Level 2: | Preferred Term - Clinical Sign Level 3: | Clinical Signs IMDRF Codes # | Preferred Term-Health Impact Level 1: | Preferred Term- Health Impact Level 2: | Preferred Term- Health Impact Level 3: | Health Impact IMDRF Codes # | Cause of Investigation-Type of Investigation IMDRF Codes # | Preferred Term - Investigation Finding Level 1: | Cause of Investigation - Investigation Finding IMDRF Codes # | Preferred Term - Investigation Finding Level 1: | Preferred Term - Type/cause of investigation Level 2: | Preferred Term - Type/cause of investigation Level 3: | Cause of Investigation- Investigation Conclusion IMDRF Codes # | Preferred Term - Investigation Conclusion Level 1: | Preferred Term - Investigation Conclusion Level 2: | Investigation Status of the AEs/AIs (Done /Not Done) | Causality Assessment (Unrelated, Possible, Probable & Certain) | Risk Assessment (Critical/High/Medium/Low) | Regulatory action(s) taken | 1st Assessor Name | Date of assessment | 2nd Assessor Name | Date of assessment | Acknowledgement / Feedback
MEDICAL DEVICE AND IN VITRO DIAGNOSTIC ADVERSE EVENTS/INCIDENTS REGISTER TMDA/DMD/MDV/R/002 Rev #: 01
1 | AEMD/2026-27/001 | 2026-06-07 | Revital Infusion giving set | Infusion giving set | NOT STATED | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | Kenya | MSD/GOT | Blockage | 2026-05-10 | 2026-05-10 | BIMA pharmacy | Dodoma | Intial | Grace Katunzi, P.O.BOX DODOMA RH 0717788929katunzigrace@gmail.com | NO | Mechanical | Hose | None | G04069 | Material intergrity problem | Burst container or vessel | Not Stated | A0402 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | CONDUCT FIELD INVESTIGATION | SUZANA T RUHOGORA | 21/7/2026
2 | AEMD/2026-27/002 | 2026-06-07 | Surgimed Medical Supplies Ltd | Insulin Syringe with needle | 1 ml | 23973711 | MD | 2023-09-01 | 2026-08-28 | Surgimed Medical Supplies Ltd |  | MSD/GOT | Blockage | 2026-06-29 | 2026-06-30 | DRRH 9094 | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Syringe | None | G04127 | Chemical problem | Device Ingridient problem | Clumpinng in device or device ingridient | A0302 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | CONDUCT FIELD INVESTIGATION AND ENHANCE MONITORING | SUZANA T RUHOGORA | 23/7/2026
3 | AEMD/2026-27/003 | 2026-06-07 | i syringe | sterile syringe | 10 ml | 251104 | MD | 2025-11-01 | 2030-10-01 | Medical Dawila Ltd. PlotP14562, Kigogo Kisarawe II, Kigamboni, Daer essalaam | TANZANIA | MSD/GOT | Blockage 1ml insulin syringe | 30/6/2026 | 2026-06-30 | EMD | Dodoma | Initial | Edina mlindasungu, DRRH DODOMA 0755359935, augustineedina@gmail.com | NO | Mechanical | syringe | None | G04127 | Chemical problem | Device Ingridient problem | Clumpinng in device or device ingridient | A0302 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | CONDUCT FIELD INVESTIGATION AND ENHANCE MONITORING | SUZANA T RUHOGORA | 23/7/2026
4 | AEMD/2026-27/004 | 2026-06-07 | Surgical face mask | Face mask |  | 260216 | MD | 2026-02-01 | 2028-01-01 | MSD P.O.BOX 9081 DAR ES SALAAM | TANZANIA | MSD | Facemask without straps and some with broken straps | 2026-06-15 | 2026-06-15 | DRRH 904 | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Mask | None | G04086 | Manufacturing ,packaging,shipping problem | Device ingridient problem | Material intergrity problem | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Request for labelling and design changes and Request Post marketing Studies | Suzana T Ruhogora | 24/7/2026
5 | AEMD/2026-27/005 | 2026-06-07 | HEUER | Infusion giving set |  |  | MD | 2026-06-01 | 2030-12-01 | Encore Medics Ltd P.O.BOX 10186, Dar essalaam | TANZANIA | MSD | Blokage |  | 2026-06-16 | BIMA pharmacy | Dodoma | Initial | Grace Katunzi, P.O.BOX DODOMA RH 0717788929katunzigrace@gmail.com | NO | Mechanical | Hose | None | G04069 | Material intergrity problem | Burst container or vessel | Not Stated | A0402 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Conduct field investigation | Suzana T Ruhogora | 24/7/2026
6 | AEMD/2026-27/006 | 2026-06-07 |  | Powdered Medical examination gloves | medium | 51042 | MD | 2025-12-01 | 2030-11-01 | Quba medical products | TANZANIA | MSD | The patient developed an itch that caused him to develop hives | 2026-12-06 |  | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | YES | Others | Appropriate term/code not available | None | G07002 | Chemical problem | Improper Chemical reaction | Not Stated | A0303 | immune system | allergic reaction | N/A | E0401 | minor injury | allergic reaction | N/A | G04127 | [10 empty] | N/A |  | Request samples for laboratory analysis and conduct field investigation | Suzana T Ruhogora | 29/7/2026
7 | AEMD/2026-27/007 | 2026-06-07 | Surgimed | Disposable insulin syringe with needle | 1ml/1cc | 23973711 | MD | Sept-2023 | 2028-08-01 | Surgimed Medical Supplies Ltd | CHINA | MSD/GOT | Blockage | 29/06/2026 | 2026-06-30 | MATERNITY PHARMACY | Dodoma | Initial | P.M DODOMA CC 079689449, Petermutal@gmail.com | NO | Mechanical | syringe | None | G004127 | Chemical problem | device ingridient problem | clumping in device or device ingridient problem | A0302 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Conduct field investigations | Suzana T Ruhogora | 31/7/2026
8 | AEMD/2026-27/008 | 2026-06-07 | Revital | Infusion giving set | N/A | 2601304-02 | MD | 2030-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD/GOT | No infusion control | 20/06/2026 | 20/06/2026 |  | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Tube | N/A | G04134 | Infusion or Flow Problem | No Flow | None | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Done | Not applicable since the incident in assessment is the adverse incident and not an event | Low | Conduct Field investigation | Gurisha Mjema | 29/07/2026
9 | AEMD/2026-27/009 | 2026-06-07 | Surgimed | AUTO DISABLE SYRINGE WITH NEEDLE | 2ml | 24985402 | MD | 2025-01-01 | 2029-02-01 | Surgimed Medical Supplies Ltd | CHINA | MSD/GOT | Blokage | 20/06/2026 | 22/06/2026 | DRRH | Dodoma | Initial | Lilian Ndabazi, DRRH 0675729301,ndabazililian@gmail.com | NO | Mechanical | Syringe | N/A | G04127 | Manufacturing, Packaging or Shipping Problem | Shipping Damage or Problem | None | A0207 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Not Done | Not applicable since the incident in assessment is the adverse incident and not an event | Low | Conduct Field investigation and Request Sample for laboratory analysis | Gurisha Mjema | 30/07/2026
10 | AEMD/2026-27/010 | 2026-06-07 | Surgimed | syringe | 2cc | 24985402 | MD | 2025-01-01 | 2029-12-01 | Surgimed Medical Supplies Ltd | CHINA | MSD/GOT | Blokage | 17/6/2026 | 16/06/2026 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO | Mechanical | Syringe | N/A | G04126 | Manufacturing, Packaging or Shipping Problem | Shipping Damage or Problem | None | A0207 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Not Done | Not applicable since the incident in assessment is the adverse incident and not an event | Low | Conduct Field investigation and Request Sample for laboratory analysis | Gurisha Mjema | 30/07/2026
11 | AEMD/2026-27/011 | 2026-06-07 | Revital | Infusion giving set |  | 2601304-02 | MD |  |  | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD/GOT | Blokage | 2026-09-06 | 2026-06-09 | DRRH | Dodoma | Initial | Furaha Ngowi,DRRH 904 ,DODMA0748132342 | NO | Mechanical | Hose | N/A | GO4069 | Material intergrity problem | Burst container or vessel | None | A0402 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Request samples for laboratory analysis and conduct field investigation | Suzana T Ruhogora | 2026-03-08
12 | AEMD/2026-27/012 | 2026-06-07 | Surgimed | AUTO DISABLE SYRINGE WITH NEEDLE | 2cc | 24985402 | MD | 2025-01-01 | 2029-12-01 | Surgimed Medical Supplies Ltd | CHINA | MSD/GOT | Blokage | 20/06/2026 | 2026-06-20 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO | Mechanical | Syringe | N/A | G04127 | Infusion or Flow Problem | None | none | A14 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NA | [10 empty] | N/A |  | Conduct field investigation and Request post marketing Studies | Suzana T Ruhogora | 2026-04-08
13 | AEMD/2026-27/013 | 2026-06-07 | Kaicare | Sterile pre-powdered syringe | 7.5 | SSJ-2025 | MD | 2025-10-01 | 2028-09-01 | VKJ health care private Ltd | INDIA |  | Holes and easy to tear | 2026-11-05 | 2026-05-11 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | YES | Others | Appropriate term/code not available | Not Stated | G00702 | Material intergrity problem | Material Perforation | material puncture /hole | A041001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Conduct field investigation | Suzana T Ruhogora | 2026-05-08
14 | AEMD/2026-27/014 | 2026-06-07 | Surgical face mask | Face mask |  | 200427 | MD | 2026-04-01 | 2028-03-01 | MSD P.O.BOX 9081 DAR ES SALAAM | TANZANIA | MSD | No ropes holding it and some ropes wasn't attached | 2026-06-06 | 2026-06-06 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO | mechanical | mask | Not Stated | G04086 | Manufacturing,shipping problem, material | Device misassembled during manufacturing | Component missing | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A | High | Conduct field investigation and Request post marketing Studies | Samwel G. Mwaijumba | 2026-05-08
15 | AEMD/2026-27/015 | 2026-06-07 | Surgical face mask | Face mask |  | 200427 | MD | 2026-04-01 | 2028-03-01 | MSD P.O.BOX 9081 DAR ES SALAAM | TANZANIA | MSD | No ropes holding it and some ropes wasn't attached | 2026-09-06 | 2026-06-09 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Mask | Not Stated | G04086 | Manufacturing ,Shipping problem,Material intergrity problem | Device Misassembled during Manufacturing/shipping,material Intergrity Problem | Component missing | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A |  | Request for labelling and design changes and request post marketing | Suzana T Ruhogora | 2026-06-08
16 | AEMD/2026-27/016 | 2026-06-07 | Kaicare | Sterile pre-powdered Gloves | 7.5 | SSJ-2025 | MD | 2025-10-01 | Sept-2028 | VKJ health care private Ltd | INDIA |  | Small size than the original size | 2026-06-09 | 2026-06-09 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | cover | non | G0437 | Manufacturing, packaging or shiping problem | product quality | not stated | A0201 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A0201 | Trend analysis | B12 | Mechanical problem identified | incorrect dimension | not stated | C0707 | Cause traced to manufacturing | quality control deficiency | D0302 | N/A | low | Conduct field investigation and Request post marketing Studies | Samwel G. Mwaijumba | 2026-11-08
17 | AEMD/2026-27/017 | 2026-06-07 | Revital | Infusion giving set |  | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD | Blokage | 2026-06-25 | 2026-07-01 | DRRH | Dodoma | Initial | Lilian Ndabazi, DRRH 0675729301,ndabazililian@gmail.com | NO | Mechanical | Tube | Not stated | GO4134 | Infusion or flow problem | No flow | not stated | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A1408 | Trend analysis | B12 | Mechanical problem identified | Brockage identified | not stated | C0708 | Cause traced to manufacturing | Manufacturing deficiency | D030 | N/A | low | Conduct field investigation and Request post marketing Studies | Samwel G. Mwaijumba | 2026-12-08
18 | AEMD/2026-27/018 | 2026-06-07 |  | AUTO DISABLE SYRINGE WITH NEEDLE | 10CC | 250904 | MD | SEPT-2025 | 2030-08-01 | Medical Dawila Ltd. PlotP14562, Kigogo Kisarawe II, Kigamboni, Daer essalaam | TANZANIA |  | Blokage | 2026-06-02 | 2026-06-02 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Syringe | Not stated | G04127 | Infusion or flow problem | No flow | not stated | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | N/A | low | Conduct field investigation and Request post marketing Studies | Samwel G. Mwaijumba | 13/08/2026
19 | AEMD/2026-27/019 | 2026-06-07 | MSD FFP2/N95 RESPIRATORY MASK | Surgical face mask |  | 260216 | MD | 2026-02-01 | 2028-01-01 | MSD P.O.BOX 9081 DAR ES SALAAM |  | MSD | No ropes holding it and some ropes wasn't attached |  |  | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Mask | Not stated | G04086 | Manufacturing, packaging or shiping problem | Device misassembled during manufacturing | component missing | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A020102 | Trend analysis | B12 | Manufacturing process problem identified | assemble problem identified | packaging incorrect or incomplete device | C160504 | Cause traced to manufacturing | quality control deficiency | D0302 | N/A | low | Conduct field investigation and Request post marketing Studies | Samwel G. Mwaijumba | 13/08/2026
20 | AEMD/2026-27/020 | 2026-06-07 | Revital | Infusion giving set |  | 260130402 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | Kenya |  | Blokage | 2026-06-01 | 2026-06-01 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO
21 | AEMD/2026-27/021 | 2026-06-07 | Revital | Infusion giving set |  | 260130402 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | Kenya | MSD | Blokage |  | 2026-05-23 | DRRH | Dodoma | Initial | Jesca Mrema, 904, DRRH, DODOMA 0713177408 | NO | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A030 | insufficient information available | B22 | No findings available | Not Stated | Not Stated | C20 | Cause Not established | N/A | D15 | N/A | N/A | No sufficient information of the adverse incident or event | Suzana T Ruhogora | 21/08/2026
22 | AEMD/2026-27/022 | 2026-06-07 |  | Surgical face mask |  | 260216 | MD | 2026-02-01 | 2028-01-01 | Medipharm Manufacture company Ltd. Dar es salaam | TANZANIA | MSD | No ropes holding it and some ropes wasn't attached | 25/05/2026 | 2026-05-25 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO
23 | AEMD/2026-27/023 | 2026-06-07 |  | sterile safety syringe | 10 cc | 251104 | MD | 2025-11-01 | 2030-10-01 | Encore Medics Ltd P.O.BOX 10186, Dar essalaam | TANZANIA | MSD | Blokage | 15/05/2026 | 2026-06-07 | DRRH | Dodoma | Initial | Anthony nicodemus ntambi DRRH, DODOMA 0622478339 | NO | Mechanical | Syringe | N/A | G04127 | Chemical problem | Device ingredient or reagent problem | Coajulation in device or device ingredient | A030202 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A010 | Testing of the device from the same lot/batch rtained by manufacturer | B03 | [7 empty] | N/A | medium | Request for samples and laboratory analysis and conduct field investigation | Alex Severine Panga | 27/08/2026
24 | AEMD/2026-27/024 | 2026-06-07 | revital | Infusion giving set |  | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | Kenya | MSD | Fluid does not flow properly | 2026-05-17 | 2026-05-17 | DRRH | Dodoma | Initial | Dyness Gabriel DRRH, DODOMA 0786020323, dynessgabriel14@gmail.com | NO | Mechanical | Tube | N/A | G04134 | Infusion or flow problem | No flow | None | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A019 | Historical data analysis | B11 | N/A | N/A | N/A | N/A | N/A | N/A | Not done |  | medium | conduct field investigation | Alex Severine Panga | 27/08/2026
25 | AEMD/2026-27/025 | 2026-06-07 | Lister | sterile surgical blade |  | MAGN25111 | MD | 2025-11-01 | 2028-11-01 | Magna marketing |  |  | blunt | 2026-05-18 | 2026-05-18 | DRRH | Dodoma | Initial | Edina mlindasungu, DRRH DODOMA 0755359935, augustineedina@gmail.com | YES | Mechanical | Cutter/blade |  | G04041 | manufacturing,packaging, shipping problem | product quality | Dull blunt | A02010 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [12 empty] | low | Request samples for laboratory analysis,conduct field investigation and monitoring | Samwel G. Mwaijumba | 18/08/2026
26 | AEMD/2026-27/026 |  | Lister | sterile surgical blade | not stated | MAGN25111 | MD | 2025-11-01 | 2028-11-01 | Magna marketing |  |  | blunt | 2026-05-10 | 2026-05-13 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | YES | Mechanical | Cutter/blade | Not stated | G04041 | manufacturing,packaging, shipping problem | product quality | Dull blunt | A02010 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [12 empty] | low | request samples for laboratory analysis, conduct field investigation and clse monitorin | Samwel G. Mwaijumba | 19/08/2026
27 | AEMD/2026-27/027 | 2026-06-07 | Medi plus | Infusion giving set |  | D-12225 | MD | 2025-07-01 | 2030-06-01 | Shree umya surgical pvt.ltd | India | Medi plus | Blokage | 2026-05-05 | 2026-05-18 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | tube | Not stated | G04131 | Infusion or flow problem | No flow | Not stated | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [12 empty] | low | Request samples for laboratory analysis and conduct field investigation | Samwel G. Mwaijumba | 20/08/2026
28 | AEMD/2026-27/028 | 2026-06-07 | MSD FFP2/N95 RESPIRATORY MASK | Surgical face mask | not stated | 260216 | MD | 2026-02-01 | 2028-01-01 | Medical Store Department | Tanzania | MSD | No ropes holding it and some ropes wasn't attached | 2026-02-06 | 2026-06-02 | DRRH | Dodoma | Initial | Edina mlindasungu, DRRH DODOMA 0755359935, augustineedina@gmail.com | NO | Mechanical | Mask | Not stated | G04086 | Manufacturing and packaging | Device ingridient problem | Material integrity | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A020102 | Trend analysis | B12 | Manufacturing, process problem identified | assembly problem identified | packaging incorrect or incomplete device | C160504 | Cause traced to manufacturing | quality control deficiency | D0302 | Unrelated | High | Request samples for laboratory analysis and conduct field investigation | Samwel G. Mwaijumba | 20/08/2026
29 | AEMD/2026-27/029 | 2026-06-07 |  | Surgical face mask | not stated | 260216 | MD | 2026-02-01 | 2028-01-01 | Medical Store Department | Tanzania | MSD | No ropes holding it and some ropes wasn't attached | 2026-02-01 | 2026-06-01 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO | Mechanical | Mask | Not stated | G04086 | Manufacturing and packaging | Device ingridient problem | Material integrity | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A020102 | Trend analysis | B12 | Manufacturing process problem identified | assembly problem identified | packaging incorrect or incomplete device | C160504 | Cause traced to manufacturing | quality control deficiency | D0302 | Unrelated | high | Request samples for laboratory analysis and conduct field investigation | Samwel G. Mwaijumba | 20/08/2026
30 | AEMD/2026-27/030 | 2026-06-07 | Neovac | blood giving set | 21G | 20251105 | MD | 2025-05-01 | 2030-04-01 | jiangsu zhengkang medical apparatus ltd | China |  | pathway blocked | 2026-05-12 | 2026-05-12 | DRRH | Dodoma | Initial | Seline guambasa , DODOMA 0755062052 | NO | Mechanical | Nozzle | None | G04093 | Infusion or flow problem | No flow | none | A1408 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A010 | Testing of the device from the same lot/batch retained by manufacturer | B02 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Low | Request samples for laboratory analysis and conduct field investigation | Alex Severine Panga | 27/08/2026
31 | AEMD/2026-27/031 | 2026-06-07 |  |  |  | 0 | MD | 2026-02-01 | 2028-01-01 | Medipharm manufacturing company ltd | TANZANIA | MSD | Matokeo ya uchunguzi yanayotolewa na kifaa hicho siyo sahihi | 20/05/2025 | 2026-06-07 | DRRH | Dodoma | Initial | Abel Karavuze, DRRH 904 DODOMA 0767054810 abelkaravuze@gmail.com | NO | Mechanical | Mask | Not stated | G04086 | Manufacturing and packaging | Device ingridient problem | material integrity | A020602 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A020102 | Trend analysis | B12 | Manufacturing,process, problem identified | assembly problem identified | packaging incorrect or incomplete device | C160504 | cause traced to manufacturing | quality control deficiency | D0302 | Unrelated | High | Manufacturer to conduct field investigation | Samwel G. Mwaijumba | 26/08/2026
32 | AEMD/2026-27/032 | 2026-06-07 | TECOM TC220 | Tecom automatic Hematology Analyzer | N/A | C9-E0430 | MD | 2026-07-01 | Not indicated | Tecom Science Corporation Jiangxi 330096 P.R CHINA | CHINA | MSD |  | 2025-12-01 | 2026-06-26 | Kamsamba Health Center | Songwe | Initial | TS, Kamsamba Health Center P.O.BOX 273 Tunduma | NO
33 | AEMD/2026-27/033 | 2026-06-07 | Clindiag SA Series | semi-automated chemistry analyzer | N/A | SASMKO47 | MD | 21/04/2021 | Not indicated | Clindiag Systems Co., Ltd Jiangsu, China | CHINA | MSD | Kifaa hakifanyi kazi | Not Indicated | 17/06/2026 | Isansa Health Center | Mbozi-Songwe | initial | BB, Isansa Health Center, P.O.BOX 70, Mbozi-Songwe | NO | Measurement | Calibrator | Not stated | G03003 | Calibrator problem | Failure to calibrate | Not stated | A0801 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Testing actual/suspend device | B01 | Operation problem identified | N/A | N/A | C13 | Cause traced to device design | Design inadequate with purpose | D0101 | Probable | Medium | An immediatly device correction is required | IBRAHIM N. SAID | 2026-09-09
34 | AEMD/2026-27/034 | 2026-06-07 | DIRUI | Chemistry Analyzer | N/A | 247000DOR0087 | MD | Sept-2024 | Not indicated | Dirui industrial Co.,Ltd Jilin 130012, P.R. China | CHINA | MSD | Kifaa kina hitilafu katika vigezo vyake | Not Indicated | 24/06/2026 | Momba Distric Hospital | Songwe | Initial | HM, Momba District Hospital, P.O.BOX 273 Tunduma | NO | Measurement | Calibrator | Not stated | G03003 | Calibrator problen | Not stated | Not stated | A0801 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Testing actual/suspend device | B01 | N/A | Failure to calibrate or used out of calibration | N/A | C1302 | Cause traced to manufacture | N/A | D03 | Probable | Medium | A Device removal for corrective action. | IBRAHIM N. SAID | 2026-09-09
35 | AEMD/2026-27/035 | 2026-06-07 | DIRUI | Chemistry Analyzer | N/A | 247000DOR0087 | MD | 2024-09-01 | Not indicated | Dirui industrial Co.,Ltd Jilin 130012, P.R. China | CHINA | MSD | Kifaa hakifanyi kazi, control unit hazikufaulu wakati wa majaribio ya kifaa hicho | Not Indicated | 19/06/2026 | Tunduma TC, P.O.BOX 73, | Songwe | Initial | Tunduma TC, P.O.BOX 73, | NO | Measurement | Calibrator | Not stated | G03003 | Calibrator problem | Not stated | Not stated | A0801 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Testing actual/suspend device | B01 | Patient sample problem | known interferent | N/A | C1402 | Cause traced to design | Design inadequate with purpose | D0101 | Probable | Medium | Device analysis ans suspension of market licence | ALBERT K. MATIKU | 2026-09-09
36 | AEMD/2026-27/036 | 2026-06-07 | YUWELL | Oxygen concetrator | N/A | 200600022 | MD | 2024-09-01 | Not indicated | Dirui industrial Co.,Ltd Jilin 130012, P.R. China | CHINA | MSD | kifaa kinazalisha O2 kwa kiwango kidogo sana | Not Indicated | 2026-03-06 | P.O.Box 01, Isoko Ileje | Songwe | initial | P.O.Box 01, Isoko Ileje | YES | Mechanical | Automatic injection system | N/A | G4007 | Output problem | Gas Output Problem | Not stated | A0906 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Testing actual/suspend device | B01 | N/A | N/A | N/A | N/A | Cause traced to design | N/A | D01 | Probable | Medium | Servicing and perfom flow calibration | IBRAHIM N. SAID | 2026-10-10
37 | AEMD/2026-27/037 | 2026-06-07 | ARI | Oxygen concetrator |  | Not stated | MD | 2025-03-01 | Not indicated | ARI medical Technology Co.Ltd | CHINA | MSD | kifaa kinazalisha O2 kwa kiwango kidogo sana | Not Indicated | 29/06/2026 | Songwe DH, P.O.Box 77 | Songwe | Initial | Z.K Songwe DH, P.O.Box 77 | YES | N/A | OXYGEN ANALYZER | N/A | G0300101 | Output problem | Gas Output Problem | Not stated | A0906 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Testin actual device | B01 | N/A | N/A | N/A | N/A | Cause traced to design | N/A | D01 | Probable | Medium | Laboratory analysis to detecr effect and suspension of device | ALBERT K. MATIKU | 2026-10-10
38 | AEMD/2026-27/038 | 2026-06-07 | ARI | Oxygen concetrator | N/A | DM20240806127 | MD | 2024-08-01 | Not indicated | ARI medical Technology Co.Ltd | CHINA | MSD | kifaa kinazalisha O2 kwa kiwango kidogo sana | 30/06/2026 | 30/06/2026 | Mbuyuni Health Center P.O.Box 77 | Songwe | initial | B.M Mbuyuni Health Center P.O.Box 77 | NO | Measurement | Aperture | Not stated | G03002 | Output problem | Gas output problem | N/A | A0906 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A19 | Historical data analysis | B11 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Medium risk | Conduct field investigationat facility to check user operation log, perform a physical inspection in the components. | Alex Severine Panga | 2026-01-09
39 | AEMD/2026-27/039 | 20/07/2026 | TRUSKIN | Sterile Disposable Latex Surgical Gloves | 7.5 | E21030601 | MD | 2021-03-03 | 2026-02-01 | Healthium meditech limited | INDIA | HEALTHIUM MEDTECH PRIVATE LIMITED | Gloves zinachanika wakati wa kuvaa during caesarean section | 15/01/2026 | 20/01/2026 | Huruma Hospital, Rombo,P.O.Box 948 | Rombo | inital | M.P.S ,0714957152,Huruma Hospital ,Rombo,P.O.Box 948 | NO | Other terms or codes related to the component | Appropriate component term or code not available | Not stated | G07002 | Material integrity problem | N/A | N/A | A04 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | A10 | Testing the device from the same lot / batch retained by manufacturer | B02 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Low | The incident was low risk however samples should be analyzed to determine root cause with continued monitoring for similar incidents. | Alex Severine Panga | 2026-01-09
41 | AEMD/2026-27/041 | 2026-05-09 |  | IV SET WITH 150ML BURETTE WITH FILTER | 150ML | 2507396 | MD | 2025-01-07 | 31/06/2030 | NUBENO HEALTHCARE PRIVATE LIMITED ,Mihan Nagpur-441108, Maharashra | INDIA | MSD | IV fluid drops are not entering the drip chamber properly and the flow regulator is malfunctioning so allowing the infusion rate to exceed the intended dose this map esults pateient to receive excessivert. | 2026-04-09 | 2026-04-09 | Katavi Regional Referral Hospital. | KATAVI | Initial | Drogo sanya, 0765240923-KATAVI
42 | AEMD/2026-27/042 | 2026-07-09 |  | AUTODISABLE SYRINGE WITHY NEEDLE | 2CC | 26993805 | MD | 2026-02-04 | 2031-03-01 | SURGMED MEDICAL SUPPLIES.SHEOXING ZHIJANG -CHINA | CHINA | MSD | Zikiingia ndani bomba halirudi | 2026-03-03 | 2026-07-09 | Tumbi RRH - Kibaha | Pwani | Initial | M.R.M, 0746778893- Kibaha | NO
43 | AEMD/2026-27/043 | 2026-07-09 | Neo gloves | Examination gloves | Medium | 2130425612P2A | MD | 24/12/2025 | 13/12/2030 | Neomedic Limited Lot 4969,Sesongr.D.E Malaysia | MALAYSIA | MSD | Medical personel started itching after wearing the gloves | 2026-10-08 | 2026-10-08 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO
44 | AEMD/2026-27/044 | 2026-07-09 | IV CANNULA WITH WINGS |  | 16G | 582502 | MD | 2025-02-01 | 2030-01-01 | Neomedic Limited,INDIA | INDIA | MSD | Empty package without cannula |  |  |  |  | Initial |  | NO
45 | AEMD/2026-27/045 | 2026-07-09 | Revital | Infusion giving test | 21G*11/2 | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD | Broken Spike IV giving set | 2026-10-08 |  | DRRH | Dodoma | Initial | Abel Karavuze,abelkaravuze@gmail.com,0767054810 | NO
46 | AEMD/2026-27/046 | 2026-07-09 | Revital | Infusion giving test | 21G*11/2 | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD | The infusion set conatined particals inside it(contaminants) | 2026-08-07 |  | DRRH | Dodoma | Initial | Grace Katunzo,Katunzograce@gmail.com,0717788929 | NO
47 | AEMD/2026-27/047 | 2026-07-09 | Revital | Infusion giving test | 21G*11/2 | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD | IV giving set had a blunt needle and had a blockage | 13/07/2026 |  | DRRH | Dodoma | Initial | Abel Karavuze,abelkaravuze@gmail.com,0767054810 | NO
48 | AEMD/2026-27/048 | 2026-07-09 | Revital | Infusion giving test | 21G*11/2 | 2601304-02 | MD | 2026-01-01 | 2030-12-01 | Revital Healthcare (EPZ) Ltd LR No.5025/1239 Takaungu 80100 Mombasa | KENYA | MSD | Absence of the regulator part of the infusion set | 24/07/2025 |  | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO
49 | AEMD/2026-27/049 | 2026-07-09 | ARI | blood pressure monitor |  |  | MD | 24/09/2022 |  | ARI medical technology co. ltd |  | MSD | kifaa hakikuweza kutumika kupima wagonjwa wa pressure | 2026-03-04 |  | Tumbi RRH - Kibaha | Pwani | Initial | M.R.M, 0746778893- Kibaha | NO
50 | AEMD/2026-27/050 | 2026-07-09 | DERIC 2000 | Non rebreather oxygen mask |  | DEC 2201360 | MD | 2022-02-01 | 2027-01-01 | DERIC 2000 Co Ltd, p.o.box 11357- Kipata mnyamwezi | CHINA | DERIC 2000 | The bag leaked and passed the air out | 2026-07-01 |  | Amana Hospitakl | Dar es salaam | Initial | Arafat Abdallah, 0693444545 | NO
51 | AEMD/2026-27/052 | 2026-07-09 | EL Dawliaico | Sterile hypodermic syringe | 60ml | 2407 | MD | 2024-07-01 | 2029-06-01 | BQ. General Wahis 53, 1030 Brussels- Belgium Mail@obelis.net | CHINA | OBELIS | The nozzel of the syringe is positioned on the periphery of the tip, thus making the pressure of the feeding contents very low compared to centrally located nozzel |  | 27/06/2026 |  |  | Initial | LUSUNGU S LUPENZA, 0765065630 | NO
52 | AEMD/2026-27/053 | 2026-08-09 | KAS Biotech Limited | HIV rapid test cassette |  | 100925lO1 | IVD | 2026-11-25 | 2026-11-27 | KAS Biotech Limited Umoja Complex, Plot No. 09, Ground Floor, Vingunguti Industrial Area, Nyerere Road, Dar es Salaam, Tanzania | Tanzania | MSD | Upon conducting lot and lot verification it was found that 5 samples from known client with positive HIV tested negative 5 samples from negative clients tested positive | 04/05/2026 to 28/05/2026 |  | Mwl. Nyerere Memorial RRH Musoma MC | MARA | initial | Kenedy Fussy 0717000880 | YES
53 | AEMD/2026-27/054 | 2026-08-09 | Panbio | COVID-19 Ag Rapid Test Device |  | 41ADJ025B | IVD | 2025-04-06 | 2027-03-06 | Abbott Rapid Diagnostic Jena Gmbh, Orlaweg 1, 07743, germany | Germany | Bahari pharmacy ltd | The incident involved a pediatric patient in the emergency department. During nasal collection, the swab shaft fractured , leaving a portion of the swab inside patient nasal cavity. The patient wassubsequently referred to operating room for swab removal . No further information on patient outcome was reported as per 12 Aug 2026 | 2026-11-08 | 2026-11-08 |  |  |  | Mary Katwaza mck@baharipharmacy.co.tz | YES
54 | AEMD/2026-27/055 | 2026-09-09 | MSD easy breath | Surgical Face Mask |  | 260427 | MD | 2026-04-01 | 2028-03-01 | MSD MEDIPHARM MANUFACTURING COMPANY LIMITED | Tanzania | MSD | face mask without belt | 2026-10-07 |  | DRRH | Dodoma | Initial | Abel Karavuze,abelkaravuze@gmail.com,0767054810 | NO
55 | AEMD/2026-27/056 | 2026-09-09 | i Syringe 10 cc | Syringe 10ml | 10cc | 251104 | MD | 2025-11-01 | 2030-10-01 | Medical Dawlia Ltd P14502,Kigamboni,Dar es salaam | Tanzania | MSD | Blockage of the syringe | 2026-07-07 | 2026-07-07 | DRRH | Dodoma | Initial | Christian christian, DRRH 240-Dodoma 065793477, christianmapunda@gmail.com | NO
56 | AEMD/2026-27/057 | 2026-09-09 | Crepe Bandage | Crepe Bandage | 7.5cm*4cm |  | MD | 2023-09-01 | 2028-08-01 |  |  | MSD | The medical device was not sufficient in size didnt have the actual size labelled |  | 24/08/2026 | DRRH | Dodoma | Initial | Edna Mlindasungu,agustoedina@gmail.com,0753359935 | NO
57 | AEMD/2026-27/058 | 2026-06-07 |  | SODIUM CHLORIDE INTRAVENOUS INFUSION BP (0.9% w/v) | 500ml | FA1401114 | MD | 2026-01-26 | 2026-12-27 | ALFA PHARMACEUTICAL LIMITED DAR ES SALAAM | TANZANIA |  | SUCSEPTIBLE NORMAL SALINE WITHOUT A LEBEL | 22/06/2026 | 23/06/2026 |  | DODOMA | initial | COSTANSIA MTWANGILI ,DODOMA | NO
57 | AEMD/2026-27/059 | 2026-10-09 | one plus infusion set | infusion set |  | 0C031 | MD | 2026-02-01 | 2031-01-01 | ilife medical devices pvt ltd | INDIA | MSD | rubber ina achia na dawa inamwagika (under normal pressure) |  | 2026-01-09 |  | DAR ES SALAAM | intial | isaya D , 0683938693 | NO
58 | AEMD/2026-27/060 | 2026-10-09 | medi plus infusion set | INFUSION SET |  | D-12225 | MD | 2025-07-01 | 2030-06-01 | Shree Umiya Surgical PVT.LTD | India | MSD | Regulator haikuweza kupitisha fluid mda wa kutumia |  | 31/08/2026 | Temeke RRH | Temeke | Initial | Ndayisaba Habona, 0718 875 356 | NO